## Supported Commands
Supported | 34461A SCPI | Notes
Configuration Commands
Supported | CONFigure?
Supported | CONFigure:CAPacitance [{¦AUTO¦MIN¦MAX¦DEF} [, {¦MIN¦MAX¦DEF}]]
Supported | CONFigure:CONTinuity
Partial Support | CONFigure:CURRent:{AC¦DC} [{¦AUTO¦MIN¦MAX¦DEF} [, {¦MIN¦MAX¦DEF}]] | Ranges different on 6500
Supported | CONFigure:DIODe
Partial Support | CONFigure:{FREQuency¦PERiod} [{¦MIN¦MAX¦DEF} [, {¦MIN¦MAX¦DEF}]] | Both parameters are ignored
Partial Support | CONFigure:{RESistance¦FRESistance} [{<range>¦AUTO¦MIN¦MAX¦DEF} [, {<resolution>¦MIN¦MAX¦DEF}]] | Ranges different on 6500
Partial Support | CONFigure:TEMPerature [{FRTD¦RTD¦FTHermistor¦THERmistor¦TCouple¦DEFault} [, {<type>¦DEFault} [,1 ,{<resolution>¦MIN¦MAX¦DEF}]]]] | TCouple Does not apply to 34461 FTHermistor not available on 6500, Type is ignored, Resolution is ignored
Partial Support | CONFigure[:VOLTage]:{AC¦DC} [{<range>¦AUTO¦MIN¦MAX¦DEF} [, {<resolution>¦MIN¦MAX¦DEF}]] | Range more limited on 6500 for AC
Partial Support | CONFigure[:VOLTage][:DC]:RATio [{<range>¦AUTO¦MIN¦MAX¦DEF} [, {<resolution>¦MIN¦MAX¦DEF}]] | Ratio shares all config variables with DCV
Measurement Commands
Supported | MEASure:CAPacitance? [{<range>¦AUTO¦MIN¦MAX¦DEF} [, {<resolution>¦MIN¦MAX¦DEF}]]
Supported | MEASure:CONTinuity?
Partial Support | MEASure:CURRent:{AC¦DC}? [{<range>¦AUTO¦MIN¦MAX¦DEF} [, {<resolution>¦MIN¦MAX¦DEF}]] | Ranges different on 6500
Supported | MEASure:DIODe?
Partial Support | MEASure:{FREQuency¦PERiod}? [{<range>¦MIN¦MAX¦DEF} [, {<resolution>¦MIN¦MAX¦DEF}]] | Both parameters are ignored
Partial Support | MEASure:{RESistance¦FRESistance}? [{<range>¦AUTO¦MIN¦MAX¦DEF} [, {<resolution>¦MIN¦MAX¦DEF}]] | Ranges different on 6500
Partial Support | MEASure:TEMPerature? [{FRTD¦RTD¦FTHermistor¦THERmistor¦TCouple¦DEFault} [, {<type>¦DEFault} [,1 ,{<resolution>¦MIN¦MAX¦DEF}]]]] | TCouple Does not apply to 34461 FTHermistor not available on 6500, Type is ignored, Resolution is ignored
Partial Support | MEASure[:VOLTage]:{AC¦DC}? [{<range>¦AUTO¦MIN¦MAX¦DEF} [, {<resolution>¦MIN¦MAX¦DEF}]] | Range more limited on 6500 for AC
Partial Support | MEASure[:VOLTage][:DC]:RATio? [{<range>¦AUTO¦MIN¦MAX¦DEF} [, {<resolution>¦MIN¦MAX¦DEF}]] | Ratio shares all config variables with DCV
Measurement Configuration Commands
Supported | [SENSe:]FUNCtion[:ON] "<function>"
Supported | [SENSe:]FUNCtion[:ON]?
AC and DC Voltage and DC Ratio Configuration Commands
Partial Support | CONFigure[:VOLTage]:{AC¦DC} [{<range>¦AUTO¦MIN¦MAX¦DEF} [, {<resolution>¦MIN¦MAX¦DEF}]] | Range more limited on 6500 for AC
Partial Support | CONFigure[:VOLTage][:DC]:RATio [{<range>¦AUTO¦MIN¦MAX¦DEF} [, {<resolution>¦MIN¦MAX¦DEF}]] | Ratio shares all config variables with DCV
Supported | CONFigure?
Supported | [SENSe:]VOLTage:AC:BANDwidth {<filter>¦MIN¦MAX¦DEF} | Options are slightly different
Supported | [SENSe:]VOLTage:AC:BANDwidth? [{MIN¦MAX¦DEF}]
Supported | [SENSe:]VOLTage:{AC¦DC}:NULL[:STATe] {OFF¦ON}
Supported | [SENSe:]VOLTage:{AC¦DC}:NULL[:STATe]?
Partial Support | [SENSe:]VOLTage:{AC¦DC}:NULL:VALue {<value>¦MIN¦MAX¦DEF} | 6500 more limited range
Supported | [SENSe:]VOLTage:{AC¦DC}:NULL:VALue? [{MIN¦MAX¦DEF}]
Supported | [SENSe:]VOLTage:{AC¦DC}:NULL:VALue:AUTO {OFF¦ON} | Will need to have an auto state, should this also turn rel on?
Supported | [SENSe:]VOLTage:{AC¦DC}:NULL:VALue:AUTO?
Partial Support | [SENSe:]VOLTage:{AC¦DC}:RANGe {<range>¦MIN¦MAX¦DEF} | Range more limited on 6500 for AC
Supported | [SENSe:]VOLTage:{AC¦DC}:RANGe? [{MIN¦MAX¦DEF}]
Supported | [SENSe:]VOLTage:{AC¦DC}:RANGe:AUTO {OFF¦ON¦ONCE}
Supported | [SENSe:]VOLTage:{AC¦DC}:RANGe:AUTO?
Not Supported | [SENSe:]VOLTage:AC:SECondary {"OFF"¦"CALCulate:DATA"¦"FREQuency"¦"VOLTage[:DC]"} | No true secondary function on 6500
Not Supported | [SENSe:]VOLTage:AC:SECondary? | No true secondary function on 6500
Not Supported | [SENSe:]VOLTage[:DC]:APERture {<seconds>¦MIN¦MAX¦DEF} | 34461 Does not have this option
Not Supported | [SENSe:]VOLTage[:DC]:APERture? [{MIN¦MAX¦DEF}] | 34461 Does not have this option
Not Supported | [SENSe:]VOLTage[:DC]:APERture:ENABled {OFF¦ON} | 34461 Does not have this option
Not Supported | [SENSe:]VOLTage[:DC]:APERture:ENABled? | 34461 Does not have this option
Not Supported | [SENSe:]VOLTage[:DC]:IMPedance:AUTO {OFF¦ON} | Not available on 6500
Not Supported | [SENSe:]VOLTage[:DC]:IMPedance:AUTO? | Not available on 6500
Partial Support | [SENSe:]VOLTage[:DC]:NPLC {<PLCs>¦MIN¦MAX¦DEF} | NPLC is more limited on 6500
Supported | [SENSe:]VOLTage[:DC]:NPLC? [{MIN¦MAX¦DEF}]
Not Supported | [SENSe:]VOLTage[:DC]:RATio:SECondary {"OFF"¦"CALCulate:DATA"¦"SENSe:DATA"} | No true secondary function on 6500
Not Supported | [SENSe:]VOLTage[:DC]:RATio:SECondary? | No true secondary function on 6500
Not Supported | [SENSe:]VOLTage[:DC]:RESolution {<resolution>¦MIN¦MAX¦DEF} | Not currently supported
Not Supported | [SENSe:]VOLTage[:DC]:RESolution? [{MIN¦MAX¦DEF}] | Not currently supported
Not Supported | [SENSe:]VOLTage[:DC]:SECondary {"OFF"¦"CALCulate:DATA"¦"VOLTage:AC"¦"PTPeak"} | No true secondary function on 6500
Not Supported | [SENSe:]VOLTage[:DC]:SECondary? | No true secondary function on 6500
Partial Support | [SENSe:]VOLTage[:DC]:ZERO:AUTO {OFF¦ON¦ONCE} | Once only works if current function matches
Supported | [SENSe:]VOLTage[:DC]:ZERO:AUTO?
2-Wire and 4-Wire Resistance Configuration Commands
Partial Support | CONFigure:{RESistance¦FRESistance} [{<range>¦AUTO¦MIN¦MAX¦DEF} [, {<resolution>¦MIN¦MAX¦DEF}]] | Ranges different on 6500
Supported | CONFigure?
Not Supported | [SENSe:]{RESistance¦FRESistance}:APERture {<seconds>¦MIN¦MAX¦DEF} | 34461 Does not have this option
Not Supported | [SENSe:]{RESistance¦FRESistance}:APERture? [{MIN¦MAX¦DEF}] | 34461 Does not have this option
Not Supported | [SENSe:]{RESistance¦FRESistance}:APERture:ENABled {OFF¦ON} | 34461 Does not have this option
Not Supported | [SENSe:]{RESistance¦FRESistance}:APERture:ENABled? | 34461 Does not have this option
Partial Support | [SENSe:]{RESistance¦FRESistance}:NPLC {<PLCs>¦MIN¦MAX¦DEF} | NPLC goes lower on 6500 but higher on 34461
Supported | [SENSe:]{RESistance¦FRESistance}:NPLC? [{MIN¦MAX¦DEF}]
Supported | [SENSe:]{RESistance¦FRESistance}:NULL[:STATe] {OFF¦ON}
Supported | [SENSe:]{RESistance¦FRESistance}:NULL[:STATe]?
Partial Support | [SENSe:]{RESistance¦FRESistance}:NULL:VALue {<value>¦MIN¦MAX¦DEF} | Null value slightly more limited on 6500
Supported | [SENSe:]{RESistance¦FRESistance}:NULL:VALue? [{MIN¦MAX¦DEF}]
Supported | [SENSe:]{RESistance¦FRESistance}:NULL:VALue:AUTO {OFF¦ON}
Supported | [SENSe:]{RESistance¦FRESistance}:NULL:VALue:AUTO?
Not Supported | [SENSe:]{RESistance¦FRESistance}:OCOMpensated {OFF¦ON} | Does not apply to 34461
Not Supported | [SENSe:]{RESistance¦FRESistance}:OCOMpensated? | Does not apply to 34461
Not Supported | [SENSe:]{RESistance¦FRESistance}:POWer:LIMit[:STATe] {OFF¦ON} | Does not apply to 34461
Not Supported | [SENSe:]{RESistance¦FRESistance}:POWer:LIMit[:STATe]? | Does not apply to 34461
Partial Support | [SENSe:]{RESistance¦FRESistance}:RANGe {<range>¦MIN¦MAX¦DEF} | Ranges go higher on 34461 but lower on 6500
Supported | [SENSe:]{RESistance¦FRESistance}:RANGe? [{MIN¦MAX¦DEF}]
Supported | [SENSe:]{RESistance¦FRESistance}:RANGe:AUTO {OFF¦ON¦ONCE}
Supported | [SENSe:]{RESistance¦FRESistance}:RANGe:AUTO?
Not Supported | [SENSe:]{RESistance¦FRESistance}:RESolution {<resolution>¦MIN¦MAX¦DEF} | Not currently supported
Not Supported | [SENSe:]{RESistance¦FRESistance}:RESolution? [{MIN¦MAX¦DEF}] | Not currently supported
Not Supported | [SENSe:]{FRESistance¦RESistance}:SECondary {"OFF"¦"CALCulate:DATA"} | No true secondary function on 6500
Not Supported | [SENSe:]{FRESistance¦RESistance}:SECondary? | No true secondary function on 6500
Partial Support | [SENSe:]RESistance:ZERO:AUTO {OFF¦ON¦ONCE} | Once only works if current function matches
Supported | [SENSe:]RESistance:ZERO:AUTO?
AC and DC Current Configuration Commands
Partial Support | CONFigure:CURRent:{AC¦DC} [{<range>¦AUTO¦MIN¦MAX¦DEF} [, {<resolution>¦MIN¦MAX¦DEF}]] | Ranges different on 6500
Supported | CONFigure?
Supported | [SENSe:]CURRent:AC:BANDwidth {<filter>¦MIN¦MAX¦DEF} | Options are slightly different
Supported | [SENSe:]CURRent:AC:BANDwidth? [{MIN¦MAX¦DEF}]
Supported | [SENSe:]CURRent:{AC¦DC}:NULL[:STATe] {OFF¦ON}
Supported | [SENSe:]CURRent:{AC¦DC}:NULL[:STATe]?
Partial Support | [SENSe:]CURRent:{AC¦DC}:NULL:VALue {<value>¦MIN¦MAX¦DEF} | Null value slightly more limited on 6500
Supported | [SENSe:]CURRent:{AC¦DC}:NULL:VALue? [{MIN¦MAX¦DEF}]
Supported | [SENSe:]CURRent:{AC¦DC}:NULL:VALue:AUTO {OFF¦ON}
Supported | [SENSe:]CURRent:{AC¦DC}:NULL:VALue:AUTO?
Partial Support | [SENSe:]CURRent:{AC¦DC}:RANGe {<range>¦MIN¦MAX¦DEF} | Ranges different on 6500
Supported | [SENSe:]CURRent:{AC¦DC}:RANGe? [{MIN¦MAX¦DEF}]
Supported | [SENSe:]CURRent:{AC¦DC}:RANGe:AUTO {OFF¦ON¦ONCE}
Supported | [SENSe:]CURRent:{AC¦DC}:RANGe:AUTO?
Not Supported | [SENSe:]CURRent:{AC¦DC}:TERMinals {3¦10} | 6500 only has 3A terminal
Supported | [SENSe:]CURRent:{AC¦DC}:TERMinals?
Not Supported | [SENSe:]CURRent:AC:SECondary {"OFF"¦"CALCulate:DATA"¦"FREQuency"¦"CURRent[:DC]"} | No true secondary function on 6500
Not Supported | [SENSe:]CURRent:AC:SECondary? | No true secondary function on 6500
Not Supported | [SENSe:]CURRent[:DC]:APERture {<seconds>¦MIN¦MAX¦DEF} | 34461 Does not have this option
Not Supported | [SENSe:]CURRent[:DC]:APERture? [{MIN¦MAX¦DEF}] | 34461 Does not have this option
Not Supported | [SENSe:]CURRent[:DC]:APERture:ENABled {OFF¦ON} | 34461 Does not have this option
Not Supported | [SENSe:]CURRent[:DC]:APERture:ENABled? | 34461 Does not have this option
Partial Support | [SENSe:]CURRent[:DC]:NPLC {<PLCs>¦MIN¦MAX¦DEF} | NPLC goes lower on 6500 but higher on 34461
Supported | [SENSe:]CURRent[:DC]:NPLC? [{MIN¦MAX¦DEF}]
Not Supported | [SENSe:]CURRent[:DC]:RESolution {<resolution>¦MIN¦MAX¦DEF} | Not currently supported
Not Supported | [SENSe:]CURRent[:DC]:RESolution? [{MIN¦MAX¦DEF}] | Not currently supported
Not Supported | [SENSe:]CURRent[:DC]:SECondary {"OFF"¦"CALCulate:DATA"¦"CURRent:AC"¦"PTPeak"} | No true secondary function on 6500
Not Supported | [SENSe:]CURRent[:DC]:SECondary? | No true secondary function on 6500
Supported | [SENSe:]CURRent[:DC]:ZERO:AUTO {OFF¦ON¦ONCE}
Supported | [SENSe:]CURRent[:DC]:ZERO:AUTO?
Not Supported | [SENSe:]CURRent:SWITch:MODE {FAST¦CONTinuous} | Does not apply to 34461
Not Supported | [SENSe:]CURRent:SWITch:MODE? | Does not apply to 34461
Capacitance Configuration Commands
Supported | CONFigure:CAPacitance [{<range>¦AUTO¦MIN¦MAX¦DEF} [, {<resolution>¦MIN¦MAX¦DEF}]]
Supported | CONFigure?
Supported | [SENSe:]CAPacitance:NULL[:STATe] {OFF¦ON}
Supported | [SENSe:]CAPacitance:NULL[:STATe]?
Supported | [SENSe:]CAPacitance:NULL:VALue {<value>¦MIN¦MAX¦DEF}
Supported | [SENSe:]CAPacitance:NULL:VALue? [{MIN¦MAX¦DEF}]
Supported | [SENSe:]CAPacitance:NULL:VALue:AUTO {OFF¦ON}
Supported | [SENSe:]CAPacitance:NULL:VALue:AUTO?
Supported | [SENSe:]CAPacitance:RANGe {<range>¦MIN¦MAX¦DEF}
Supported | [SENSe:]CAPacitance:RANGe? [{MIN¦MAX¦DEF}]
Supported | [SENSe:]CAPacitance:RANGe:AUTO {OFF¦ON¦ONCE}
Supported | [SENSe:]CAPacitance:RANGe:AUTO?
Not Supported | [SENSe:]CAPacitance:SECondary{"OFF"¦"CALCulate:DATA"} | No true secondary function on 6500
Not Supported | [SENSe:]CAPacitance:SECondary? | No true secondary function on 6500
Temperature Configuration Commands
Partial Support | CONFigure:TEMPerature [{FRTD¦RTD¦FTHermistor¦THERmistor¦TCouple¦DEFault} [, {<type>¦DEFault} [,1 ,{<resolution>¦MIN¦MAX¦DEF}]]]] | TCouple Does not apply to 34461 FTHermistor not available on 6500, Type is ignored, Resolution is ignored
Supported | CONFigure?
Not Supported | [SENSe:]TEMPerature:APERture {<seconds>¦MIN¦MAX¦DEF} | 34461 Does not have this option
Not Supported | [SENSe:]TEMPerature:APERture? [{MIN¦MAX¦DEF}] | 34461 Does not have this option
Not Supported | [SENSe:]TEMPerature:APERture:ENABled {OFF¦ON} | 34461 Does not have this option
Not Supported | [SENSe:]TEMPerature:APERture:ENABled? | 34461 Does not have this option
Partial Support | [SENSe:]TEMPerature:NPLC {<PLCs>¦MIN¦MAX¦DEF} | NPLC goes lower on 6500 but higher on 34461
Supported | [SENSe:]TEMPerature:NPLC? [{MIN¦MAX¦DEF}]
Supported | [SENSe:]TEMPerature:NULL[:STATe] {OFF¦ON}
Supported | [SENSe:]TEMPerature:NULL[:STATe]?
Partial Support | [SENSe:]TEMPerature:NULL:VALue {<value>¦MIN¦MAX¦DEF} | Values limited by 6500
Partial Support | [SENSe:]TEMPerature:NULL:VALue? [{MIN¦MAX¦DEF}] | Values limited by 6500
Supported | [SENSe:]TEMPerature:NULL:VALue:AUTO {OFF¦ON}
Supported | [SENSe:]TEMPerature:NULL:VALue:AUTO?
Not Supported | [SENSe:]TEMPerature:SECondary {"OFF"¦"CALCulate:DATA"¦"SENSe:DATA"} | No true secondary function on 6500
Not Supported | [SENSe:]TEMPerature:SECondary? | No true secondary function on 6500
Not Supported | [SENSe:]TEMPerature:TRANsducer:{FRTD¦RTD}:OCOMpensated {OFF¦ON} | Does not apply to 34461
Not Supported | [SENSe:]TEMPerature:TRANsducer:{FRTD¦RTD}:OCOMpensated? | Does not apply to 34461
Not Supported | [SENSe:]TEMPerature:TRANsducer:{FRTD¦RTD}:POWer:LIMit[:STATe] {OFF¦ON} | Does not apply to 34461
Not Supported | [SENSe:]TEMPerature:TRANsducer:{FRTD¦RTD}:POWer:LIMit[:STATe]? | Does not apply to 34461
Supported | [SENSe:]TEMPerature:TRANsducer:{FRTD¦RTD}:RESistance[:REFerence] {<reference>¦MIN¦MAX¦DEF}
Supported | [SENSe:]TEMPerature:TRANsducer:{FRTD¦RTD}:RESistance[:REFerence]? [{MIN¦MAX¦DEF}]
Not Supported | [SENSe:]TEMPerature:TRANsducer:{FTHermistor¦THERmistor}:POWer:LIMit[:STATe] {OFF¦ON} | Does not apply to 34461
Not Supported | [SENSe:]TEMPerature:TRANsducer:{FTHermistor¦THERmistor}:POWer:LIMit[:STATe]? | Does not apply to 34461
Not Supported | [SENSe:]TEMPerature:TRANsducer:TCouple:CHECk {OFF¦ON} | Does not apply to 34461
Not Supported | [SENSe:]TEMPerature:TRANsducer:TCouple:CHECk? | Does not apply to 34461
Not Supported | [SENSe:]TEMPerature:TRANsducer:TCouple:RJUNction {<temperature>¦MIN¦MAX¦DEF} | Does not apply to 34461
Not Supported | [SENSe:]TEMPerature:TRANsducer:TCouple:RJUNction? [{MIN¦MAX¦DEF}] | Does not apply to 34461
Not Supported | [SENSe:]TEMPerature:TRANsducer:TCouple:RJUNction:OFFSet:ADJust {<temperature>¦MIN¦MAX¦DEF} | Does not apply to 34461
Not Supported | [SENSe:]TEMPerature:TRANsducer:TCouple:RJUNction:OFFSet:ADJust? [{MIN¦MAX¦DEF}] | Does not apply to 34461
Not Supported | [SENSe:]TEMPerature:TRANsducer:TCouple:RJUNction:TYPE {INTernal¦FIXed} | Does not apply to 34461
Not Supported | [SENSe:]TEMPerature:TRANsducer:TCouple:RJUNction:TYPE? | Does not apply to 34461
Not Supported | [SENSe:]TEMPerature:TRANsducer:TCouple:TYPE {E¦J¦K¦N¦R¦T} | Does not apply to 34461
Not Supported | [SENSe:]TEMPerature:TRANsducer:TCouple:TYPE? | Does not apply to 34461
Partial Support | [SENSe:]TEMPerature:TRANsducer:TYPE {FRTD¦RTD¦FTHermistor¦THERmistor¦TCouple} | TCouple Does not apply to 34461 FTHermistor not available on 6500
Partial Support | [SENSe:]TEMPerature:TRANsducer:TYPE? | TCouple Does not apply to 34461 FTHermistor not available on 6500
Supported | [SENSe:]TEMPerature:ZERO:AUTO {OFF¦ON¦ONCE}
Supported | [SENSe:]TEMPerature:ZERO:AUTO?
Supported | UNIT:TEMPerature {C¦F¦K}
Supported | UNIT:TEMPerature?
Frequency/Period Configuration Commands
Partial Support | CONFigure:{FREQuency¦PERiod} [{<range>¦MIN¦MAX¦DEF} [, {<resolution>¦MIN¦MAX¦DEF}]] | Both parameters are ignored
Supported | CONFigure?
Not Supported | [SENSe:]{FREQuency¦PERiod}:APERture {<seconds>¦MIN¦MAX¦DEF} | 34461 Does not have this option
Not Supported | [SENSe:]{FREQuency¦PERiod}:APERture? [{MIN¦MAX¦DEF}] | 34461 Does not have this option
Supported | [SENSe:]{FREQuency¦PERiod}:NULL[:STATe] {OFF¦ON}
Supported | [SENSe:]{FREQuency¦PERiod}:NULL[:STATe]?
Partial Support | [SENSe:]{FREQuency¦PERiod}:NULL:VALue {<value>¦MIN¦MAX¦DEF} | Null value slightly more limited on 6500
Partial Support | [SENSe:]{FREQuency¦PERiod}:NULL:VALue? [{MIN¦MAX¦DEF}] | Null value slightly more limited on 6500
Supported | [SENSe:]{FREQuency¦PERiod}:NULL:VALue:AUTO {OFF¦ON}
Supported | [SENSe:]{FREQuency¦PERiod}:NULL:VALue:AUTO?
Not Supported | [SENSe:]{FREQuency¦PERiod}:RANGe:LOWer {<freq>¦MIN¦MAX¦DEF} | Freq/Per range not supported on 6500
Not Supported | [SENSe:]{FREQuency¦PERiod}:RANGe:LOWer? | Freq/Per range not supported on 6500
Not Supported | [SENSe:]{FREQuency¦PERiod}:TIMeout:AUTO {OFF¦ON} | No timeout control on 6500
Not Supported | [SENSe:]{FREQuency¦PERiod}:VOLTage:RANGe {<range>¦MIN¦MAX¦DEF} | Freq/Per range not supported on 6500
Not Supported | [SENSe:]{FREQuency¦PERiod}:VOLTage:RANGe? [{MIN¦MAX¦DEF}] | Freq/Per range not supported on 6500
Not Supported | [SENSe:]{FREQuency¦PERiod}:VOLTage:RANGe:AUTO {OFF¦ON¦ONCE} | Freq/Per range not supported on 6500
Not Supported | [SENSe:]{FREQuency¦PERiod}:VOLTage:RANGe:AUTO? | Freq/Per range not supported on 6500
Not Supported | [SENSe:]FREQuency:SECondary {"OFF"¦"CALCulate:DATA"¦"PERiod"¦"VOLTage:AC"} | No true secondary function on 6500
Not Supported | [SENSe:]FREQuency:SECondary? | No true secondary function on 6500
Not Supported | [SENSe:]PERiod:SECondary {"OFF"¦"CALCulate:DATA"¦"FREQuency"¦"VOLTage:AC"} | No true secondary function on 6500
Not Supported | [SENSe:]PERiod:SECondary? | No true secondary function on 6500
Continuity and Diode Configuration Commands
Supported | CONFigure:CONTinuity
Supported | CONFigure:DIODe
Secondary Measurement Commands
Not Supported | [SENSe:]CAPacitance:SECondary {"OFF"¦"CALCulate:DATA"} | No true secondary function on 6500
Not Supported | [SENSe:]CAPacitance:SECondary? | No true secondary function on 6500
Not Supported | [SENSe:]CURRent:AC:SECondary {"OFF"¦"CALCulate:DATA"¦"FREQuency"¦"CURRent[:DC]"} | No true secondary function on 6500
Not Supported | [SENSe:]CURRent:AC:SECondary? | No true secondary function on 6500
Not Supported | [SENSe:]CURRent[:DC]:SECondary {"OFF"¦"CALCulate:DATA"¦"CURRent:AC"¦"PTPeak"} | No true secondary function on 6500
Not Supported | [SENSe:]CURRent[:DC]:SECondary? | No true secondary function on 6500
Not Supported | [SENSe:]DATA2? | No true secondary function on 6500
Not Supported | [SENSe:]DATA2:CLEar[:IMMediate] | No true secondary function on 6500
Not Supported | [SENSe:]{FRESistance¦RESistance}:SECondary {"OFF"¦"CALCulate:DATA"} | No true secondary function on 6500
Not Supported | [SENSe:]{FRESistance¦RESistance}:SECondary? | No true secondary function on 6500
Not Supported | [SENSe:]FREQuency:SECondary {"OFF"¦"CALCulate:DATA"¦"PERiod"¦"VOLTage:AC"} | No true secondary function on 6500
Not Supported | [SENSe:]PERiod:SECondary {"OFF"¦"CALCulate:DATA"¦"FREQuency"¦"VOLTage:AC"} | No true secondary function on 6500
Not Supported | [SENSe:]TEMPerature:SECondary {"OFF"¦"CALCulate:DATA"¦"SENSe:DATA"} | No true secondary function on 6500
Not Supported | [SENSe:]TEMPerature:SECondary? | No true secondary function on 6500
Not Supported | [SENSe:]VOLTage:AC:SECondary {"OFF"¦"CALCulate:DATA"¦"FREQuency"¦"VOLTage[:DC]"} | No true secondary function on 6500
Not Supported | [SENSe:]VOLTage:AC:SECondary? | No true secondary function on 6500
Not Supported | [SENSe:]VOLTage[:DC]:RATio:SECondary {"OFF"¦"CALCulate:DATA"¦"SENSe:DATA"} | No true secondary function on 6500
Not Supported | [SENSe:]VOLTage[:DC]:RATio:SECondary? | No true secondary function on 6500
Not Supported | [SENSe:]VOLTage[:DC]:SECondary {"OFF"¦"CALCulate:DATA"¦"VOLTage:AC"¦"PTPeak"} | No true secondary function on 6500
Not Supported | [SENSe:]VOLTage[:DC]:SECondary? | No true secondary function on 6500
Miscellaneous Configuration Commands
Supported | ROUTe:TERMinals?
Sample Commands
Supported | SAMPle:COUNt {<count>¦MIN¦MAX¦DEF}
Supported | SAMPle:COUNt? [{MIN¦MAX¦DEF}]
Not Supported | SAMPle:COUNt:PRETrigger {<count>¦MIN¦MAX¦DEF} | Not available on 34461A
Not Supported | SAMPle:COUNt:PRETrigger? [{MIN¦MAX¦DEF}] | Not available on 34461A
Not Supported | SAMPle:SOURce {IMMediate¦TIMer} | Not available on 34461A
Not Supported | SAMPle:SOURce? | Not available on 34461A
Not Supported | SAMPle:TIMer {<interval>¦MIN¦MAX¦DEF} | Not available on 34461A
Not Supported | SAMPle:TIMer? [{MIN¦MAX¦DEF}] | Not available on 34461A
Triggering Commands
Partial Support | ABORt | Only breaks measurements, not commands that are waiting
Supported | INITiate[:IMMediate]
Not Supported | OUTPut:TRIGger:SLOPe {POSitive¦NEGative} | No meas complete trigger to map external out to
Not Supported | OUTPut:TRIGger:SLOPe? | No meas complete trigger to map external out to
Supported | READ?
Supported | SAMPle:COUNt {<count>¦MIN¦MAX¦DEF}
Supported | SAMPle:COUNt? [{MIN¦MAX¦DEF}]
Not Supported | SAMPle:COUNt:PRETrigger {<count>¦MIN¦MAX¦DEF} | Not available on 34461A
Not Supported | SAMPle:COUNt:PRETrigger? [{MIN¦MAX¦DEF}] | Not available on 34461A
Not Supported | *TRG | Will need backdoor commands to implement
Supported | TRIGger:COUNt {<count>¦MIN¦MAX¦DEF¦INFinity}
Supported | TRIGger:COUNt? [{MIN¦MAX¦DEF}]
Supported | TRIGger:DELay {<seconds>¦MIN¦MAX¦DEF}
Supported | TRIGger:DELay? [{MIN¦MAX¦DEF}]
Supported | TRIGger:DELay:AUTO {OFF¦ON}
Supported | TRIGger:DELay:AUTO?
Not Supported | TRIGger:LEVel {<level>¦MIN¦MAX¦DEF} | Not available on 34461A
Not Supported | TRIGger:LEVel? [{MIN¦MAX¦DEF}] | Not available on 34461A
Supported | TRIGger:SLOPe {POSitive¦NEGative}
Supported | TRIGger:SLOPe?
Partial Support | TRIGger:SOURce {IMMediate¦EXTernal¦BUS¦INTernal} | BUS not working yet
Supported | TRIGger:SOURce?
Overall
Supported | CALCulate:CLEar[:IMMediate]
Histogram
Not Supported | CALCulate:TRANsform:HISTogram:ALL? | 6500 Does not support histogram commands
Not Supported | CALCulate:TRANsform:HISTogram:CLEar[:IMMediate] | 6500 Does not support histogram commands
Not Supported | CALCulate:TRANsform:HISTogram:COUNt? | 6500 Does not support histogram commands
Not Supported | CALCulate:TRANsform:HISTogram:DATA? | 6500 Does not support histogram commands
Not Supported | CALCulate:TRANsform:HISTogram:POINts {<value>¦MIN¦MAX¦DEF} | 6500 Does not support histogram commands
Not Supported | CALCulate:TRANsform:HISTogram:POINts? [{MIN¦MAX¦DEF}] | 6500 Does not support histogram commands
Not Supported | CALCulate:TRANsform:HISTogram:RANGe:AUTO {OFF¦ON} | 6500 Does not support histogram commands
Not Supported | CALCulate:TRANsform:HISTogram:RANGe:AUTO? | 6500 Does not support histogram commands
Not Supported | CALCulate:TRANsform:HISTogram:RANGe:{LOWer¦UPPer} {<value>¦MIN¦MAX¦DEF} | 6500 Does not support histogram commands
Not Supported | CALCulate:TRANsform:HISTogram:RANGe:{LOWer¦UPPer}? [{MIN¦MAX¦DEF}] | 6500 Does not support histogram commands
Not Supported | CALCulate:TRANsform:HISTogram[:STATe] {OFF¦ON} | 6500 Does not support histogram commands
Not Supported | CALCulate:TRANsform:HISTogram[:STATe]? | 6500 Does not support histogram commands
Limit Testing
Supported | CALCulate:LIMit:CLEar[:IMMediate]
Partial Support | CALCulate:LIMit:{LOWer¦UPPer}[:DATA] {<value>¦MIN¦MAX¦DEF} | Limit max/min more limited on 6500
Partial Support | CALCulate:LIMit:{LOWer¦UPPer}[:DATA]? [{MIN¦MAX¦DEF}] | Limit max/min more limited on 6500
Supported | CALCulate:LIMit[:STATe] {OFF¦ON}
Supported | CALCulate:LIMit[:STATe]?
Scaling
Supported | CALCulate:SCALe:DB:REFerence {<reference>¦MIN¦MAX¦DEF}
Supported | CALCulate:SCALe:DB:REFerence? [{MIN¦MAX¦DEF}]
Supported | CALCulate:SCALe:DBM:REFerence {<reference>¦MIN¦MAX¦DEF}
Supported | CALCulate:SCALe:DBM:REFerence? [{MIN¦MAX¦DEF}]
Supported | CALCulate:SCALe:FUNCtion {DB¦DBM¦PCT¦SCALe} | PCT and SCALe don't apply to 34461A
Supported | CALCulate:SCALe:FUNCtion?
Not Supported | CALCulate:SCALe:GAIN {<gain>¦MIN¦MAX¦DEF} | Does not apply to 34461
Not Supported | CALCulate:SCALe:GAIN? [{MIN¦MAX¦DEF}] | Does not apply to 34461
Not Supported | CALCulate:SCALe:OFFSet {<offset>¦MIN¦MAX¦DEF} | Does not apply to 34461
Not Supported | CALCulate:SCALe:OFFSet? [{MIN¦MAX¦DEF}] | Does not apply to 34461
Not Supported | CALCulate:SCALe:REFerence:AUTO {OFF¦ON} | 6500 does not have an auto reference option
Not Supported | CALCulate:SCALe:REFerence:AUTO? | 6500 does not have an auto reference option
Not Supported | CALCulate:SCALe:REFerence {<reference>¦MIN¦MAX¦DEF} | Does not apply to 34461
Not Supported | CALCulate:SCALe:REFerence? [{MIN¦MAX¦DEF}] | Does not apply to 34461
Supported | CALCulate:SCALe[:STATe] {OFF¦ON}
Supported | CALCulate:SCALe[:STATe]?
Not Supported | CALCulate:SCALe:UNIT <quoted_string> | Does not apply to 34461
Not Supported | CALCulate:SCALe:UNIT? | Does not apply to 34461
Not Supported | CALCulate:SCALe:UNIT:STATe {OFF¦ON} | Does not apply to 34461
Not Supported | CALCulate:SCALe:UNIT:STATe? | Does not apply to 34461
Statistics
Supported | CALCulate:AVERage:ALL?
Supported | CALCulate:AVERage:AVERage?
Supported | CALCulate:AVERage:CLEar[:IMMediate]
Supported | CALCulate:AVERage:COUNt?
Supported | CALCulate:AVERage:MAXimum?
Supported | CALCulate:AVERage:MINimum?
Supported | CALCulate:AVERage:PTPeak?
Supported | CALCulate:AVERage:SDEViation?
Supported | CALCulate:AVERage[:STATe] {OFF¦ON}
Supported | CALCulate:AVERage[:STATe]?
Smoothing
Not Supported | CALCulate:SMOothing:RESPonse {SLOW¦MEDium¦FAST} | Does not apply to 34461
Not Supported | CALCulate:SMOothing:RESPonse? | Does not apply to 34461
Not Supported | CALCulate:SMOothing[:STATe] {OFF¦ON} | Does not apply to 34461
Not Supported | CALCulate:SMOothing[:STATe]? | Does not apply to 34461
Trend Chart
Not Supported | CALCulate:TCHart[:STATe] {OFF¦ON} | Always on on 6500
Not Supported | CALCulate:TCHart[:STATe]? | Always on on 6500
Reading Memory Commands
Supported | DATA:LAST? | Should be able to execute at any time
Supported | DATA:POINts:EVENt:THReshold <count>
Supported | DATA:POINts:EVENt:THReshold?
Supported | DATA:POINts?
Partial Support | DATA:REMove? <num_readings> [,WAIT] | If buffer doesn't have enough readings and measurement is idle, just returns with no error
Supported | FETCh?
Supported | R? [<max_readings>]
Calibration Commands
Not Supported | *CAL? | No calibration control on 6500
Not Supported | CALibration:ADC? | No calibration control on 6500
Not Supported | CALibration[:ALL]? | No calibration control on 6500
Not Supported | CALibration:COUNt? | No calibration control on 6500
Not Supported | CALibration:DATE? | No calibration control on 6500
Not Supported | CALibration:SECure:CODE <new_code> | No calibration control on 6500
Not Supported | CALibration:SECure:STATe {OFF¦ON} [, <code>] | No calibration control on 6500
Not Supported | CALibration:SECure:STATe? | No calibration control on 6500
Not Supported | CALibration:STORe | No calibration control on 6500
Not Supported | CALibration:STRing "<string>" | No calibration control on 6500
Not Supported | CALibration:STRing? | No calibration control on 6500
Not Supported | CALibration:TEMPerature? | No calibration control on 6500
Not Supported | CALibration:TIME? | No calibration control on 6500
Not Supported | CALibration:VALue <value> | No calibration control on 6500
Not Supported | CALibration:VALue? | No calibration control on 6500
Not Supported | SYSTem:ACALibration:DATE? | No calibration control on 6500
Not Supported | SYSTem:ACALibration:TEMPerature? | No calibration control on 6500
Not Supported | SYSTem:ACALibration:TIME? | No calibration control on 6500
State Storage and Preferences Commands
Not Supported | *LRN? | Possible but time consumuming to support
Not Supported | MMEMory:LOAD:PREFerences <file> | May be able to get this working but would need to be very custom
Supported | MMEMory:LOAD:STATe <file> | This is now MMEMory:LOAD:STATe <file>[, AUTO] where if auto is specified the script name will be autoexec and run startup script, the filename parameter will be ignored, all path data is ignored
Not Supported | MMEMory:STORe:PREFerences <file> | May be able to get this working but would need to be very custom
Supported | MMEMory:STORe:STATe <file> | This is now MMEMory:STORe:STATe <file>[, AUTO] where if auto is specified the script name will be autoexec and run on startup, the filename parameter will be ignored, all path data is ignored
Not Supported | MMEMory:STATe:RECall:AUTO {OFF¦ON} | Auto is handled in STORE:STATE command
Not Supported | MMEMory:STATe:RECall:AUTO? | Auto is handled in STORE:STATE command
Not Supported | MMEMory:STATe:RECall:SELect <file> | Auto is handled in STORE:STATE command
Not Supported | MMEMory:STATe:RECall:SELect? | Auto is handled in STORE:STATE command
Supported | MMEMory:STATe:VALid? <file> | This is now MMEMory:STATe:VALid <file>[, AUTO] where if auto is specified the script name will be autoexec and check startup script, the filename parameter will be ignored, all path data is ignored
Supported | *RCL {0¦1¦2¦3¦4}
Supported | *SAV {0¦1¦2¦3¦4}
General Purpose File Management Commands
Partial Support | MMEMory:CATalog[:ALL]? [<folder>[<filespec>]] | Simply lists all file names in a specified directory in block data format, only USB availbale
Supported | MMEMory:CDIRectory <folder> | Default directory not reset between instances of program, only USB availbale
Supported | MMEMory:CDIRectory? | Default directory not reset between instances of program, only USB availbale
Not Supported | MMEMory:COPY <file1>, <file2> | No Ccopy function on 6500
Not Supported | MMEMory:DELete {<file>¦<filespec>} | Cannot delete files on 6500
Supported | MMEMory:MDIRectory <folder> | Only USB availbale
Not Supported | MMEMory:MOVE <file1>, <file2> | Cannot change filename on 6500
Supported | MMEMory:RDIRectory <folder> | Only USB availbale
Data Transfer Commands
Not Supported | MMEMory:DOWNload:DATA <binary_block> | No downloading on 6500
Not Supported | MMEMory:DOWNload:FNAMe <file> | No downloading on 6500
Not Supported | MMEMory:DOWNload:FNAMe? | No downloading on 6500
Not Supported | MMEMory:FORMat:READing:CSEParator {COMMa¦SEMicolon¦TAB} | Not supported on 34461A
Not Supported | MMEMory:FORMat:READing:CSEParator? | Not supported on 34461A
Not Supported | MMEMory:FORMat:READing:INFormation {OFF¦ON} | Not supported on 34461A
Not Supported | MMEMory:FORMat:READing:INFormation? | Not supported on 34461A
Not Supported | MMEMory:FORMat:READing:RLIMit {OFF¦ON} | Not supported on 34461A
Not Supported | MMEMory:FORMat:READing:RLIMit? | Not supported on 34461A
Supported | MMEMory:STORe:DATA RDG_STORE, <file> | Only CSV allowed
Not Supported | MMEMory:UPLoad? <file>+A379:D396 | No uploading on 6500
IEEE-488 Commands
Not Supported | *CAL? | No calibration control on 6500
Supported | *CLS
Supported | *ESE <enable_value>
Supported | *ESE?
Supported | *ESR?
Supported | *IDN?
Not Supported | *LRN? | Possible but time consumuming to support
Supported | *OPC
Supported | *OPC?
Not Supported | *OPT? | Could just return a dummy string
Not Supported | *PSC {0¦1} | Cannot save enable variables in non-volitile
Partial Support | *PSC?
Supported | *RCL {0¦1¦2¦3¦4}
Supported | *RST
Supported | *SAV {0¦1¦2¦3¦4}
Supported | *SRE <enable_value>
Supported | *SRE?
Supported | *STB?
Not Supported | *TRG | Will need backdoor commands to implement
Not Supported | *TST? | No self test on 6500
Supported | *WAI | Just works like waitcomplete()
Format Subsystem
Not Supported | FORMat:BORDer {NORMal¦SWAPped} | Does not apply to 34461
Not Supported | FORMat:BORDer? | Does not apply to 34461
Not Supported | FORMat[:DATA] {ASCii¦REAL} [, <length>] | Does not apply to 34461
Not Supported | FORMat[:DATA]? | Does not apply to 34461
System-Related Commands
Not Supported | *CAL? | No calibration control on 6500
Supported | DISPlay[:STATe] {OFF¦ON}
Supported | DISPlay[:STATe]?
Supported | DISPlay:TEXT:CLEar
Supported | DISPlay:TEXT[:DATA] "<string>"
Supported | DISPlay:TEXT[:DATA]?
Partial Support | DISPlay:VIEW {NUMeric¦HISTogram¦TCHart¦METer} | Meter,Histogram not supported on 6500
Supported | DISPlay:VIEW?
Not Supported | HCOPy:SDUMp:DATA:FORMat {PNG¦BMP} | Screenshot capability unavailable over TSP
Not Supported | HCOPy:SDUMp:DATA:FORMat? | Screenshot capability unavailable over TSP
Not Supported | HCOPy:SDUMp:DATA? | Screenshot capability unavailable over TSP
Supported | *IDN?
Not Supported | LXI:IDENtify[:STATe] {OFF¦ON} | LXI settings not accessible with TSP
Not Supported | LXI:IDENtify[:STATe]? | LXI settings not accessible with TSP
Not Supported | LXI:MDNS:ENABle {OFF¦ON} | LXI settings not accessible with TSP
Not Supported | LXI:MDNS:ENABle? | LXI settings not accessible with TSP
Not Supported | LXI:MDNS:HNAMe[:RESolved]? | LXI settings not accessible with TSP
Not Supported | LXI:MDNS:SNAMe:DESired "<name>" | LXI settings not accessible with TSP
Not Supported | LXI:MDNS:SNAMe:DESired? | LXI settings not accessible with TSP
Not Supported | LXI:MDNS:SNAMe[:RESolved]? | LXI settings not accessible with TSP
Not Supported | LXI:RESet | LXI settings not accessible with TSP
Not Supported | LXI:RESTart | LXI settings not accessible with TSP
Supported | *RST
Not Supported | SYSTem:ACALibration:DATE? | No Acal commands on 6500
Not Supported | SYSTem:ACALibration:TEMPerature? | No Acal commands on 6500
Not Supported | SYSTem:ACALibration:TIME? | No Acal commands on 6500
Supported | SYSTem:BEEPer[:IMMediate]
Partial Support | SYSTem:BEEPer:STATe {OFF¦ON} | Cant change beeper on cont or diode remotely, this is NOT non-volatile
Supported | SYSTem:BEEPer:STATe?
Not Supported | SYSTem:CLICk:STATe {OFF¦ON} | No click option on 6500 (Front panel only)
Not Supported | SYSTem:CLICk:STATe? | No click option on 6500 (Front panel only)
Supported | SYSTem:DATE <year>, <month>, <day>
Supported | SYSTem:DATE?
Supported | SYSTem:ERRor[:NEXT]?
Supported | SYSTem:HELP? | Will need to fill this in if I have time
Supported | SYSTem:IDENtify {DEFault¦AT34460A¦AT34461A¦AT34410A¦AT34411A¦HP34401A}*
Supported | SYSTem:IDENtify?
Supported | SYSTem:LABel "<string>" | This is NOT non-volatile
Supported | SYSTem:LABel?
Supported | SYSTem:LFRequency?
Partial Support | SYSTem:LOCal | This is how you exit the emulation program
Partial Support | SYSTem:PRESet | No histogram support
Not Supported | SYSTem:SECurity:COUNt? | No Security on 6500
Not Supported | SYSTem:SECurity:IMMediate | No Security on 6500
Not Supported | SYSTem:TEMPerature? | No temperature access on 6500
Supported | SYSTem:TIME <hour>, <minute>, <second>
Partial Support | SYSTem:TIME? | No partial seconds
Not Supported | SYSTem:UPTime? | No way to access this on 6500
Not Supported | SYSTem:USB:HOST:ENABle {OFF¦ON} | Cannot enable/disable front USB on 6500
Not Supported | SYSTem:USB:HOST:ENABle? | Cannot enable/disable front USB on 6500
Partial Support | SYSTem:VERSion? | Just prints 1994.0
Not Supported | SYSTem:WMESsage "<string>" | No way to set a poweron message on 6500
Not Supported | SYSTem:WMESsage? | No way to set a poweron message on 6500
Not Supported | TEST:ALL? | Could accept and just return 0
Not Supported | *TST? | Could accept and just return 0
Interface Locking Commands
Not Supported | SYSTem:LOCK:NAME? | No locking on 6500
Not Supported | SYSTem:LOCK:OWNer? | No locking on 6500
Not Supported | SYSTem:LOCK:RELease | No locking on 6500
Not Supported | SYSTem:LOCK:REQuest? | No locking on 6500
License Management Commands
Not Supported | SYSTem:LICense:CATalog? | No license on 6500
Not Supported | SYSTem:LICense:DELete "<option_name>" | No license on 6500
Not Supported | SYSTem:LICense:DELete:ALL | No license on 6500
Not Supported | SYSTem:LICense:DESCription? "<option_name>" | No license on 6500
Not Supported | SYSTem:LICense:ERRor? | No license on 6500
Not Supported | SYSTem:LICense:ERRor:COUNt? | No license on 6500
Not Supported | SYSTem:LICense:INSTall [{<folder>¦<file>}] | No license on 6500
Not Supported | SYSTem:LICense:INSTall? "<option_name>" | No license on 6500
Interface Configuration Commands
Not Supported | SYSTem:COMMunicate:ENABle {OFF¦ON}, {GPIB¦HISLip¦USB¦LAN¦SOCKets¦TELNet¦VXI11¦WEB¦USBMtp¦USBHost} | Cannot control which protocol is enabled on 6500
Not Supported | SYSTem:COMMunicate:ENABle? {GPIB¦HISLip¦USB¦LAN¦SOCKets¦TELNet¦VXI11¦WEB¦USBMtp¦USBHost} | Cannot control which protocol is enabled on 6500
Supported | SYSTem:COMMunicate:GPIB:ADDRess <address>
Supported | SYSTem:COMMunicate:GPIB:ADDRess?
Not Supported | SYSTem:COMMunicate:LAN:CONTrol? | Cannot access port number
Partial Support | SYSTem:COMMunicate:LAN:DHCP {OFF¦ON} | Just changes between auto/manual
Partial Support | SYSTem:COMMunicate:LAN:DHCP? | Just indicates auto or manual connection
Not Supported | SYSTem:COMMunicate:LAN:DNS[{1¦2}] "<address>" | Cannot set DNS on 6500
Not Supported | SYSTem:COMMunicate:LAN:DNS[{1¦2}]? [{CURRent¦STATic}] | Cannot set DNS on 6500
Not Supported | SYSTem:COMMunicate:LAN:DOMain? | No domain control on 6500
Partial Support | SYSTem:COMMunicate:LAN:GATeway "<address>" | Only sets if LAN is manual
Partial Support | SYSTem:COMMunicate:LAN:GATeway? [{CURRent¦STATic}] | Only prints current
Not Supported | SYSTem:COMMunicate:LAN:HOSTname "<name>" | No hostname control on 6500
Not Supported | SYSTem:COMMunicate:LAN:HOSTname? [{CURRent¦STATic}] | No hostname control on 6500
Partial Support | SYSTem:COMMunicate:LAN:IPADdress "<address>" | Only sets if LAN is manual
Partial Support | SYSTem:COMMunicate:LAN:IPADdress? [{CURRent¦STATic}] | Only prints current
Supported | SYSTem:COMMunicate:LAN:MAC?
Partial Support | SYSTem:COMMunicate:LAN:SMASk "<mask>" | Only sets if LAN is manual
Partial Support | SYSTem:COMMunicate:LAN:SMASk? [{CURRent¦STATic}] | Only prints current
Not Supported | SYSTem:COMMunicate:LAN:TELNet:PROMpt "<string>" | Cannot change prompts on 6500
Partial Support | SYSTem:COMMunicate:LAN:TELNet:PROMpt? | Just returns "TSP>"
Not Supported | SYSTem:COMMunicate:LAN:TELNet:WMESsage "<string>" | No welcome messages on 6500
Not Supported | SYSTem:COMMunicate:LAN:TELNet:WMESsage? | No welcome messages on 6500
Not Supported | SYSTem:COMMunicate:LAN:UPDate | Not needed, using ipconfig()
Not Supported | SYSTem:COMMunicate:LAN:WINS[{1¦2}] "<address>" | Not available on 6500
Not Supported | SYSTem:COMMunicate:LAN:WINS[{1¦2}]? [{CURRent¦STATic}] | Not available on 6500
Not Supported | SYSTem:USB:HOST:ENABle {OFF¦ON} | Cannot enable/disable front USB on 6500
Not Supported | SYSTem:USB:HOST:ENABle? | Cannot enable/disable front USB on 6500
Status System Commands
Supported | *CLS
Supported | *ESE <enable_value>
Supported | *ESE?
Supported | *ESR?
Not Supported | *PSC {0¦1} | Cannot save enable variables in non-volitile
Partial Support | *PSC? | Cannot save enable variables in non-volitile
Supported | *SRE <enable_value>
Supported | *SRE?
Supported | STATus:OPERation:CONDition?
Supported | STATus:OPERation:ENABle <enable_value>
Supported | STATus:OPERation:ENABle?
Supported | STATus:OPERation[:EVENt]?
Supported | STATus:PRESet
Supported | STATus:QUEStionable:CONDition?
Supported | STATus:QUEStionable:ENABle <enable_value>
Supported | STATus:QUEStionable:ENABle?
Supported | STATus:QUEStionable[:EVENt]?
Supported | *STB?

## Status Model
Supported | 34461A SCPI | Notes
Status Byte Register
Partial Support | Error Queue | Only checks 6500 Errors
Supported | Questionable Summary
Supported | Output Buffer
Supported | Standard Event Summary
Supported | Master Summary
Supported | Standard Operation Summary
Standard Event Register
Supported | Operation Complete
Supported | Query Error
Supported | Device Error
Supported | Execution Error
Supported | Command Error
Partial Support | Power On | Couldn't get this to flip but this bit is being read from 6500
Questionable Data Register
Supported | Voltage Overload
Supported | Current Overload
Supported | Temperature Overload
Supported | Frequency Overload
Not Supported | Calibration Corrupt | No way to check for corrupt calibration on 6500
Supported | Resistance Overload
Supported | Capacitance Overload
Supported | Lower Limit Fail
Supported | Upper Limit Fail
Partial Support | Reading Memory Overflow | Does not account for removed data points
Standard Operation Register
Not Supported | Calibrating | No way to check that instrument is calibrating on 6500
Supported | Measuring
Partial Support | Waiting For Trig | Only works when trigger model has been waiting for over 100ms
Supported | Config Change | Config Change is whenever a change resets the buffer
Partial Support | Reading Memory Threshold | Does not account for removed data points
Not Supported | Instrument Locked | No instrument locking on 6500
Supported | Global Error